I.  General information:

1.  Mission name:  Senex 2013

2.  Instrument name:  Low Turbulence Inlet (LTI) and White-Light Optical Particle Counter (WLOPC)

3.  What is measured:  Sample flow; concentration and size of large (>1 µm) aerosol particles

4.  Short description of measurement technique:

The LTI, located at the FD station, decelerates a 38 lpm sample flow without inducing turbulence.  This flow is provided to instruments from the FD station to STA 4.  Turbulent boundary layer flow from the LTI is exhausted through the BF venturi.  The WLOPC is installed at the outboard FD station and in the B-cabinet.  It counts and sizes particles by measuring the intensity of scattered light sourced by a 3 W white-light LED.

5. Contact information for all personnel going to the field with this instrument:

(for multiple investigators, please list the PI or primary contact person first)

| Name | Email | Office phone | Cell phone |
| --- | --- | --- | --- |
| 1. Charles Brock | Charles.A.Brock@noaa.gov | 303-497-3795 | 303-887-2523 |
| 2. |  |  |  |
| 3. |  |  |  |
| 4. |  |  |  |
| 5. |  |  |  |
| 6. |  |  |  |

II.  Specific information:

1. Total installed weight:	73 lbs LTI, 78 lbs WLOPC

(rack, gas cylinders, hoses, cabling, pumps, inlets, permeation tubes, etc.)

| Rack weight and balance info | Allowed | Actual |
| --- | --- | --- |
| Weight, lbs.: |  |  |
| Overturning moment, in-lbs.: |  |  |

Pod weight and CG:

2.  Individual subassembly info (weights should sum to total listed above)

| Component name | Location name and flight station | Weight, lbs |
| --- | --- | --- |
| 1. LTI control box | FD rack, FS 309 | 16 |
| 2. LTI window plate, probe | FD window, FS 320 | 22 |
| 3. LTI plumbing, impactor | FD seat, FS 347 | 35 |
| 4. WLOPC sensor and plate | FD seat, FS347 | 29 |
| 5. WLOPC tubing and cabling | B2 cab, FS 378 | 5 |
| 6. WLOPC control box | B3 cab, FS 416 | 39 |
| 7. WLOPC mounting frame | B3 cab, FS 416 | 5 |

151

3.  Component power consumption in Amps

| Component name | Location name | 400 Hz 3Ø | 400 Hz 1Ø | 60 Hz | 28VDC | 28VDC WOW |
| --- | --- | --- | --- | --- | --- | --- |
| 1. LTI | FD |  |  |  | 1.2 |  |
| 2. WLOPC | B3 cab |  | 3.5 | 1.0 |  |  |
| 3. |  |  |  |  |  |  |
| 4. |  |  |  |  |  |  |
| 5. |  |  |  |  |  |  |
| 6. |  |  |  |  |  |  |
| 7. |  |  |  |  |  |  |
| 8. |  |  |  |  |  |  |
| 9. |  |  |  |  |  |  |
| 10. |  |  |  |  |  |  |
| 11. |  |  |  |  |  |  |
| 12. |  |  |  |  |  |  |
|  | Totals: |  | 3.5 | 1.0 | 1.2 |  |
|  |  | 400 Hz 3Ø | 400 Hz 1Ø | 60 Hz | 28VDC | 28VDC WOW |

4.  Inlet and exhaust information:

| Inlet/exhaust name | Location name and flight station | Hole size through hull, inches |
| --- | --- | --- |
| 1. Inlet-LTI | FD window – FS 309 | 2.5 |
| 2. BF Venturi | LIPF – FS 780 | 3 |
| 3. |  |  |
| 4. |  |  |
| 5. |  |  |
| 6. |  |  |

5.  Source of flow (name and location of pump or venturi)

| Pump name | Location name and flight station |
| --- | --- |
| 1. WLOPC linear pump | B3 cab (built in to control box) |
| 2. BF Venturi | LIPF-FS780 |
| 3. |  |
| 4. |  |
| 5. |  |
| 6. |  |

6.  Installed hazardous materials or equipment:

(only for items installed in the aircraft for use during flight)

A.  Lasers: NONE

B.  RF transmitters:  NONE

C.  Radioactive materials:  NONE

D.  Compressed gases: NONE

E.  Chemicals (solids and liquids):

| Chemical number: | 1 Silica gel |  |
| --- | --- | --- |
| Concentration | pure solid |  |
| Amount | 100 g |  |
| Container description | Fabric reinforced Teflon tube |  |
| Purpose | Dessicant |  |
| Solution pH | N/A |  |
| Spill kit provided? | No |  |

F.  Cryogens: NONE

G.  UPS and battery installation: NONE

H.  Motors

Description:  WLOPC linear pump, Gast SPP-25EBS

Motor current draw  <2A 125V 60Hz startup, 1A running

Thermal interlock enabled?  no

I.  Operator seat requests -

Test flights:  0

Transit flights:  0

Science flights:  0

7.  Data and plumbing drops

Network (Cat. 5/6 ethernet) drops requested:  1 at  B3 cab

Serial drops requested:

IRIG-B drops (BNC coax connector) requested:

Vacuum/exhaust/ emergency dump lines:  LTI exhaust to BF venturi

| Name | Flow rate, slpm | Line pressure, Torr | Pump type | Trace gas concentration(s) |
| --- | --- | --- | --- | --- |
| 1. WLOPC | <5 | ambient | linear | none |
| 2. LTI exhaust | 500 | ambient | BF Venturi | none |
| 3. |  |  |  |  |
| 4. |  |  |  |  |

Ground gas service lines (number, location, type of service):

Other gas lines (number, location, type of service):

Will you be sending data to the AOC data station?  If so, please provide the following information:

| Parameter name | Voltage range | Unit conversion |
| --- | --- | --- |
| 1. |  |  |
| 2. |  |  |
| 3. |  |  |

8.  Aircraft access

a.  flight days:

Pre-flight time requested at aircraft (hours):  1.0

Routine pre-flight ground support required?   None

Routine post-flight time requested at aircraft (hours): 0.5

Routine post-flight ground support required?  None

b.  non-flight days:

Cleaning of LTI inlet tip approximately every 5 flights. Need workstand to reach FD window.

9.  Aircraft maneuvers

No valid data acquired during in-cloud sampling due to droplet shattering on inlet. Inlet can ice up in rime/mixed/clear icing conditions.

10.  Miscellaneous

1.  Hazmat for preflight/postflight calibrations:

i)  <500 ml n-butanol for CPC

ii)  <20 mCi Po-210 in DMA

iii)  <2 liters isopropanol in particle generator

iv)  compressed N2 in size 50 cylinder

v)  nonhazardous <500g:  ammonium  sulfate, dioctyl sebacate, polystyrene latex

2.  Fabrication and sheet metal support:  None

3.  Ferry flight/check flight procedures.  No special procedures--instruments 	may be flown cold.

III.  Ground laboratory space

1.  Tampa space requests:

Requested on NMASS worksheet

2.  Field space requests:

Requested on NMASS worksheet