I.  General information:

1.  Mission name:  Senex 2013

2.  Instrument name:  Nucleation-mode aerosol size spectrometer (NMASS)

3.  What is measured:  Concentration and size of very small (<0.1 µm) aerosol particles

4.  Short description of measurement technique:

An inert vapor is condensed onto the small aerosol particles, growing them into droplets.  The droplets are counted by enclosed laser diode optical detectors.  The droplets and uncondensed vapor are trapped and the remaining flow exhausted out of the instrument.

5. Contact information for all personnel going to the field with this instrument:

(for multiple investigators, please list the PI or primary contact person first)

| Name | Email | Office phone | Cell phone |
| --- | --- | --- | --- |
| 1. Charles Brock | Charles.A.Brock@noaa.gov | 303-497-3795 | 303-887-2523 |
| 2. |  |  |  |
| 3. |  |  |  |
| 4. |  |  |  |
| 5. |  |  |  |
| 6. |  |  |  |

II.  Specific information:

1. Total installed weight:	162.5 lbs

(rack, gas cylinders, hoses, cabling, pumps, inlets, permeation tubes, etc.)

2.  Individual subassembly info (weights should sum to total listed above)

| Component name | Location name and flight station | Weight, lbs |
| --- | --- | --- |
| 1. NMASS box & accessories | AMPS pod RWS 83 | 162.5 |
| See AMPS Pod CG spreadsheet |  |  |

3.  Component power consumption in Amps

| Component name | Location name | 400 Hz 3Ø | 400 Hz 1Ø | 60 Hz | 28VDC | 28VDC WOW |
| --- | --- | --- | --- | --- | --- | --- |
| 1. NMASS Instrument | AMPS Pod |  |  |  | 15 |  |
| 2. |  |  |  |  |  |  |
| 3. |  |  |  |  |  |  |
| 4. |  |  |  |  |  |  |
|  | Totals: |  |  |  | 15 |  |
|  |  | 400 Hz 3Ø | 400 Hz 1Ø | 60 Hz | 28VDC | 28VDC WOW |

4.  Inlet and exhaust information:

| Inlet/exhaust name | Location name and flight station | Hole size through hull, inches |
| --- | --- | --- |
| 1. NMASS pod inlet | AMPS pod, starboard lower side skin | 2 |
| 2. NMASS exhaust | AMPS pod, aft lower skin | .625 |
| 3. Cooling flow | AMPS pod, midline aft port side | 3 (2x) |
| 4. |  |  |
| 5. |  |  |
| 6. |  |  |

Note that cooling flow is regulated by butterfly valves with power failure modes of CLOSED.

5.  Source of flow (name and location of pump or venturi)

| Pump name | Location name and flight station |
| --- | --- |
| 1. NMASS IDP3 scroll pump | AMPS pod |
| 2. |  |
| 3. |  |
| 4. |  |
| 5. |  |
| 6. |  |

6.  Installed hazardous materials or equipment:

(only for items installed in the aircraft for use during flight)

A.  Lasers

Type:  diode

Class:  IIIb

Wavelength:  783 nm

Output power:  40 mW

Eye-safe?  No

Beam fully contained within instrument during normal operation?  Yes.  Beam fully enclosed during testing and alignment--no safety procedures needed.

B.  RF transmitters:   NONE

C.  Radioactive materials: NONE

D.  Compressed gases:  NONE

E.  Chemicals (solids and liquids):

| Chemical number: | 1 Perfluorotributylamine (Fluorinert FC-43) |  |
| --- | --- | --- |
| Concentration | pure |  |
| Amount | 100ml |  |
| Container description | aluminum housing with buna o-rings |  |
| Purpose | inert condensing fluid, nontoxic, nonflammable |  |
| Solution pH | N/A |  |
| Spill kit provided? | No. Nontoxic, nonflammable, nonconductive, nonreactive |  |

F.  Cryogens: NONE

G.  UPS and battery installation:  NONE

H.  Motors

Description:  NMASS scroll pump motor, brushless DC

Motor current draw  28VDC, 10A for <500 ms at startup, <6A running

Thermal interlock enabled?  Yes

I. Operator seat requests -

Test flights:  0

Transit flights:  0

Science flights:  0

7.  Data and plumbing drops

Network (Cat. 5/6 ethernet) drops requested: 1 to AMPS distribution box

Serial drops requested:

IRIG-B drops (BNC coax connector) requested:

Vacuum/exhaust/ emergency dump lines:

| Name | Flow rate, slpm | Line pressure, Torr | Pump type | Trace gas concentration(s) |
| --- | --- | --- | --- | --- |
| 1. NMASS exhaust | <1.7 | ambient | scroll | Perfluorotributylamine (Fluorinert FC-43), <2000 ppmv |
| 2. |  |  |  |  |
| 3. |  |  |  |  |
| 4. |  |  |  |  |

Ground gas service lines (number, location, type of service):

Other gas lines (number, location, type of service):

Will you be sending data to the AOC data station? None.

8.  Aircraft access

a.  flight days:

Pre-flight time requested at aircraft (hours):  1.0

Routine pre-flight ground support required?

(stands, ladders, forklifts, covers, external equipment, etc.)

Routine post-flight time requested at aircraft (hours): 0.5

Routine post-flight ground support required?

(stands, ladders, forklifts, covers, external equipment, etc.)

PI will provide pod cover and cart with ground support equipment.

b.  non-flight days:

Routine external access to inlets or zenith mounts required?

(please describe location, how often, for how long, type of ground support equipment needed, weather constraints, etc.)

Please note there is zero access and zero power to the aircraft (including pods) on hard-down days.  These occur at least once every seven calendar days while in the field.

9.  Aircraft maneuvers

Briefly describe in-flight calibration frequency, duration, altitudes desired:

Briefly describe instrument sensitivity to flight conditions:

(issues during roll/pitch, ascent/descent, sampling in cloud, icing etc.)

No valid data acquired during in-cloud sampling due to droplet shattering on inlet

10.  Miscellaneous

1.  Hazmat for preflight/postflight calibrations:  None onboard aircraft. Cart with Po-210 (<20 mCi), N2 cylinder, isopropanol (<1 liter), butanol (<1 liter) underwing at AMPS pod.

2.  Fabrication and sheet metal support:  None.

3.  Ferry flight/check flight procedures.  No special procedures--instrument may be flown cold.

III.  Ground laboratory space

1.  Tampa space requests:

Power requirements:  120VAC 60 Hz, 6A

Special requests:  During integration, would like to operate the instrument with grounded wall power while instrument is installed in the AMPS pod.

Work space: 1 table (25 ft2), 2 chairs, internet access. Hangar floor space (100 ft2) for supplies, tools, carts.

2.  Field space requests:

Work space, ft2:  100 ft2 for NMASS+UHSAS+WLOPC+cloud physics probes

Number of tables/chairs: 4 tables, 4 chairs

Power requirements:  120VAC 60 Hz, 10A in work space area

Storage space, ft2:  100 ft2

Other requests: